# Algocyte MHS New Device Calibration and Quality Control Report

This report documents the HGB calibration and subsequent FBC quality control assessment for one new Algocyte MHS device. Complete all applicable fields, retain the supporting records, and use the approved criteria to determine the final device disposition.

# 1 Study Information

| Test protocol reference | TP-007-034 |
| --- | --- |
| HGB calibration protocol reference | TP-007-029 [revision] |
| Test report reference | [TR-007-034-XX] |
| Device ID | [Enter device ID] |
| Device model | [Enter model] |
| Serial number | [Enter serial number] |
| Calibration date | [DD-MMM-YYYY] |
| QC date | [DD-MMM-YYYY] |
| Laboratory location | [Enter location] |
| Operator | [Enter name] |
| Reviewer | [Enter name] |
| Reference machine and ID | [Sysmex XN-350 / enter instrument ID] |
| Report preparation date | [DD-MMM-YYYY] |

## 1.1 Device Under Evaluation

| Device | Device ID | Device model | Serial number | Release configuration |
| --- | --- | --- | --- | --- |
| [Enter unit name] | [Enter] | [Enter] | [Enter] | [Enter configuration ID] |

## 1.2 Phase Execution Summary

| Phase | Test type | Run identifier | Target batches | Valid batches | Unique samples | Repeats, overlap and exclusions |
| --- | --- | --- | --- | --- | --- | --- |
| Calibration | HGB only | [Enter] | 50 | [Enter] | [Enter] | [Enter counts and cross-reference] |
| Quality control | FBC | [Enter] | 50 | [Enter] | [Enter] | [Enter counts and cross-reference] |

# 2 References

•	TP-007-034 Algocyte MHS New Device Calibration and Analytical QC Protocol [revision]

•	TP-007-029 Algocyte MHS Analytical Calibration Protocol [revision]

•	[Controlled APT criteria document ID and revision]

•	[Controlled data analysis method ID and revision]

•	[Raw data and analysis output filenames or controlled locations]

# 3 Definitions

| Term | Definition |
| --- | --- |
| Batch ID | Unique identifier for an MHS acquisition or test record. |
| Sample ID | Identifier of the blood sample linked to a batch. |
| Device ID | Unique identifier of the Algocyte MHS device under evaluation. |
| Calibration phase | Fifty valid HGB-only batches used to derive and review the device-specific HGB calibration. |
| QC phase | Fifty valid FBC batches processed with the locked calibration and pipeline configuration for release assessment. |
| MAPD | Mean absolute percentage difference, when this is the approved acceptance metric. |
| Valid paired result | An evaluable MHS result and corresponding reference result that meet the pre-specified inclusion rules. |
| Not evaluable | An outcome used when the required data, metric or approved criterion is unavailable or cannot be validly assessed. |

# 4 Deviations and Exclusions

## 4.1 Deviations From Protocol

State whether deviations occurred. Describe each deviation, its impact assessment and approval reference. If none occurred, enter None.

| Deviation ID | Phase | Description | Impact assessment | Approval or disposition |
| --- | --- | --- | --- | --- |
| [Enter / None] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Enter] | [Enter] | [Enter] | [Enter] |

## 4.2 Excluded Results and Replacement Batches

Summarize exclusions here and provide the complete record in Appendix C. Do not exclude a result solely because it disagrees with the reference result.

| Phase | Excluded batches | Affected analytes | Primary reason | Replacement batches | Appendix reference |
| --- | --- | --- | --- | --- | --- |
| Calibration | [Enter] | HGB | [Enter / None] | [Enter] | Appendix C |
| Quality control | [Enter] | [Enter / None] | [Enter / None] | [Enter] | Appendix C |

# 5 Observations

[Enter relevant observations about samples, device behaviour, environmental conditions, data completeness or execution. Enter None if there are no observations.]

# 6 Experimental Conditions

## 6.1 Hardware Software and Pipeline

| Component | Identifier | Version | Configuration or notes |
| --- | --- | --- | --- |
| Algocyte MHS | [Device ID] | Firmware [enter] | [Enter] |
| Application | [Enter] | [Enter app version] | [Enter] |
| Processing pipeline | [Enter pipeline ID] | [Enter version] | [Enter commit, build or release reference] |
| Calibration parameter set | [Enter parameter ID] | [Enter version] | [Enter lock or approval record] |
| Reference instrument | [Instrument ID] | [Enter software version] | [Enter calibration or QC status reference] |

## 6.2 Samples and Batches

The phase-specific batch and sample identifiers are listed in Appendices A and B. Record the summary counts below, including repeated samples and any overlap between phases.

| Phase | Acquired batches | Valid batches | Unique samples | Repeated batches | Cross-phase overlap |
| --- | --- | --- | --- | --- | --- |
| HGB calibration | [Enter] | [Enter] | [Enter] | [Enter] | [Enter sample IDs / None] |
| FBC QC | [Enter] | [Enter] | [Enter] | [Enter] | [Enter sample IDs / None] |

## 6.3 Calibration and QC Configuration

| Record | Identifier and version | Review or lock evidence |
| --- | --- | --- |
| Calibration dataset | [Enter] | [Enter record] |
| Calibration record | [Enter] | [Enter approval] |
| Final HGB parameter set | [Enter] | [Enter lock evidence] |
| QC dataset | [Enter] | [Enter finalization record] |
| Pipeline configuration | [Enter] | [Enter configuration verification] |

# 7 Method of Data Analysis

## 7.1 Acceptance Metric

Controlled criteria document: [Enter document ID and revision]. Approved acceptance metric and units: [Enter]. Approved handling of zero or non-evaluable reference values: [Enter].

When the approved metric is the mean absolute percentage difference, calculate each valid pair as 100 x |MHS_i - Reference_i| / Reference_i and calculate MAPD as the mean across the n valid evaluable QC pairs.

If the controlled APT limit specifies a different statistic, use that statistic and identify it consistently in the results table.

## 7.2 Complementary Metrics

Report signed mean percentage bias, regression slope, regression intercept and Pearson correlation r for each analyte, when estimable. State the regression method and any weighting. Include scatter and difference plots and report the observed reference concentration range.

| Analysis item | Pre-specified method or source |
| --- | --- |
| Acceptance metric | [Enter definition and controlled source] |
| Regression method | [Enter method and weighting] |
| Outlier or exclusion method | [Enter controlled source] |
| Software and code version | [Enter] |
| Analysis output record | [Enter filename or controlled location] |

# 8 Results

## 8.1 HGB Calibration Summary

| Item | Result | Supporting record or explanation |
| --- | --- | --- |
| Valid HGB calibration batches | [Enter; target 50] | Appendix A |
| Unique calibration samples | [Enter] | Appendix A |
| Observed HGB reference range | [Enter range and units] | [Enter source] |
| Calibration parameter set | [Enter ID and version] | [Enter calibration record] |
| Calibration assessment | [PASS / FAIL / NOT EVALUABLE] | [Enter explanation] |
| Pipeline application verified | [Yes / No] | [Enter verification record] |

## 8.2 FBC QC Acceptance Results

Acceptance metric: [Enter approved name and definition]. Controlled APT criteria: [Enter document ID and revision]. Record the actual valid n for each analyte and explain every FAIL or NOT EVALUABLE result.

| Analyte | Result type and unit | Observed reference range | Valid n | Acceptance metric result | APT limit percent | Outcome | Explanation or exception record |
| --- | --- | --- | --- | --- | --- | --- | --- |
| HGB | [Enter] | [Enter] | [n] | [Result] | 6.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| HCT | [Enter] | [Enter] | [n] | [Result] | 9.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| RBC | [Enter] | [Enter] | [n] | [Result] | 8.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| WBC | [Enter] | [Enter] | [n] | [Result] | 15.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| NEUT | [Enter] | [Enter] | [n] | [Result] | 12.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| LYMPH | [Enter] | [Enter] | [n] | [Result] | 15.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| MXD | [Enter] | [Enter] | [n] | [Result] | 30.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| PLT | [Enter] | [Enter] | [n] | [Result] | 20.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| MCV | [Enter] | [Enter] | [n] | [Result] | 7.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| MCH | [Enter] | [Enter] | [n] | [Result] | 10.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| MCHC | [Enter] | [Enter] | [n] | [Result] | 7.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |
| RDW | [Enter] | [Enter] | [n] | [Result] | 15.0 | [PASS / FAIL / NOT EVAL.] | [Explain or reference] |

The APT limits above are carried forward from TP-007-034. Confirm the controlled source, revision and metric definition before approval of this report.

## 8.3 Complementary QC Metrics

| Analyte | Signed mean bias percent | Slope | Intercept | Pearson r | Regression method and weighting | Plot or notes |
| --- | --- | --- | --- | --- | --- | --- |
| HGB | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| HCT | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| RBC | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| WBC | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| NEUT | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| LYMPH | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| MXD | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| PLT | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| MCV | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| MCH | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| MCHC | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |
| RDW | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter plot ID or note] |

Complementary metrics are descriptive unless analyte-specific limits and decision rules were approved before execution.

## 8.4 Figures

Record the analyte-specific plot identifier in Section 8.3. Insert the approved plots after this section or provide their controlled file references below. Each plot should identify the device, analyte, units, valid n and analysis version.

| Figure package | Controlled file or location | Review note |
| --- | --- | --- |
| Scatter plots | [Enter] | [Enter / None] |
| Difference plots | [Enter] | [Enter / None] |
| Other figures | [Enter / None] | [Enter / None] |

# 9 Conclusion and Final Device Disposition

Algocyte MHS device [Device ID] completed [n] valid HGB-only calibration batches under calibration record [ID and version]. Calibration parameter set [ID and version] was applied through pipeline [version]. The device was subsequently assessed using [n] valid FBC QC batches paired with Sysmex XN-350 results.

The acceptance metric [name and definition] was compared with the analyte-specific limits in [controlled criteria document and revision]. The overall QC outcome and device disposition are recorded below.

| Calibration outcome | [PASS / FAIL / NOT EVALUABLE] |
| --- | --- |
| Overall FBC QC outcome | [PASS / FAIL / NOT EVALUABLE] |
| Final device disposition | [RELEASED / REJECTED - NOT RELEASED / NOT RELEASED - INVESTIGATION OR RETEST REQUIRED] |
| Exception applied | [No / Yes - enter approved exception record] |
| Decision rationale | [Enter concise rationale, failed or non-evaluable analytes, and required actions] |
| Retest or recalibration history | [Enter record references / None] |

# 10 Review and Approval

Approval confirms review of the calibration record, QC results, deviations, exclusions, final configuration and device disposition.

| Role | Name | Signature or controlled approval record | Date |
| --- | --- | --- | --- |
| Report prepared by | [Enter] | [Enter] | [DD-MMM-YYYY] |
| Data analysis reviewed by | [Enter] | [Enter] | [DD-MMM-YYYY] |
| Laboratory reviewer | [Enter] | [Enter] | [DD-MMM-YYYY] |
| Final approver | [Enter] | [Enter] | [DD-MMM-YYYY] |

# 11 Appendices

## Appendix A HGB Calibration Phase Batch Traceability

Device ID: [Enter device ID]. Phase: Calibration. Test type: HGB only. Add replacement rows if the phase contains more than 50 acquired batches.

| No. | Batch ID | Sample ID | Test ID | Reference result ID | Repeat or overlap | Included | Exclusion or notes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 2 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 3 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 4 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 5 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 6 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 7 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 8 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 9 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 10 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 11 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 12 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 13 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 14 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 15 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 16 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 17 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 18 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 19 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 20 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 21 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 22 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 23 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 24 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 25 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 26 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 27 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 28 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 29 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 30 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 31 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 32 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 33 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 34 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 35 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 36 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 37 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 38 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 39 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 40 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 41 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 42 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 43 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 44 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 45 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 46 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 47 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 48 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 49 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 50 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |

## Appendix B FBC QC Phase Batch Traceability

Device ID: [Enter device ID]. Phase: Quality control. Test type: FBC. Add replacement rows if the phase contains more than 50 acquired batches.

| No. | Batch ID | Sample ID | Test ID | Reference result ID | Repeat or overlap | Included | Exclusion or notes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 2 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 3 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 4 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 5 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 6 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 7 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 8 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 9 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 10 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 11 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 12 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 13 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 14 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 15 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 16 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 17 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 18 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 19 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 20 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 21 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 22 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 23 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 24 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 25 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 26 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 27 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 28 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 29 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 30 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 31 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 32 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 33 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 34 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 35 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 36 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 37 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 38 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 39 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 40 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 41 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 42 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 43 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 44 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 45 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 46 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 47 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 48 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 49 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |
| 50 | [Enter] | [Enter] | [Enter] | [Enter] | [No / describe] | [Yes / No] | [Enter if applicable] |

## Appendix C Excluded Results and Replacement Record

| No. | Phase | Batch ID | Sample ID | Analyte | Exclusion reason | Replacement batch | Approval or disposition |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 2 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 3 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 4 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 5 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 6 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 7 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 8 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 9 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 10 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 11 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 12 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 13 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 14 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 15 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 16 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 17 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 18 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 19 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |
| 20 | [Enter] | [Enter] | [Enter] | [Enter] | [Enter] | [Enter / None] | [Enter] |

## Appendix D Deviation Retest and Configuration Change History

| Record ID | Date | Phase | Event or change | Impact on validity | Action and approval reference |
| --- | --- | --- | --- | --- | --- |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |
| [Enter / None] | [Date] | [Enter] | [Enter] | [Enter] | [Enter] |